@PARAFILTR ON =

@Z_TOC TITLE = Abstract

This report describes the design and construction of a specialized radio timing receiver for the LORAN-C radionavigation system. The computer-controlled receiver provides a precision frequency source of 5 MHz and submultiples, together with a precision timing source of 1-pulse/second synchronized to Coordinated Universal Time UTC(LORAN). It is intended as a laboratory monitor and comparator for precision timekeeping equipment, but also has applications as a precision clock and timing source for computer networks and in other applications requiring precision timekeeping with respect to UTC.

Keywords: precision time and time interval, LORAN-C, timing receiver, disciplined oscillator.

Sponsored by: Defense Advanced Research Projects Agency under NASA Ames Research Center contract number NAG 2-638 and National Science Foundation grant number NCR-89-13623.

@PAGE =

@Z_TOC TITLE = Table of Contents

@Z_TOC LVL 2 = 1.	 Introduction	1

@Z_TOC LVL 2 = 2.	 LORAN-C System Operation	3

@Z_TOC LVL 3 = 2.1.	 Signalling Characteristics	4

@Z_TOC LVL 3 = 2.2.	 Establishing the Reference Epoch	4

@Z_TOC LVL 2 = 3.	 System Design Description	6

@Z_TOC LVL 3 = 3.1.	 Receiver and Preamplifier	6

@Z_TOC LVL 3 = 3.2.	 Receiver Controller	8

@Z_TOC LVL 3 = 3.3.	 System Timing and Sequencing	10

@Z_TOC LVL 3 = 3.4.	 Analog System Simulation	12

@Z_TOC LVL 2 = 4.	 Software Operations	14

@Z_TOC LVL 3 = 4.1.	 Mode 1: Receiver Offset Calibration	14

@Z_TOC LVL 3 = 4.2.	 Mode 2: Receiver Gain Calibration	14

@Z_TOC LVL 3 = 4.3.	 Mode 3: Signal Level Calibration	15

@Z_TOC LVL 3 = 4.4.	 Mode 4: Pulse-Group Search	15

@Z_TOC LVL 3 = 4.5.	 Mode 5: Cycle Search	16

@Z_TOC LVL 3 = 4.6.	 Mode 6: Phase Tracking	17

@Z_TOC LVL 3 = 4.7.	 Auxiliary Functions	17

@Z_TOC LVL 2 = 5.	 Performance Analysis	19

@Z_TOC LVL 3 = 5.1.	 Small-Signal Analysis	19

@Z_TOC LVL 3 = 5.2.	 Large-Signal Analysis	20

@Z_TOC LVL 3 = 5.3.	 The Cycle Identification Problem	21

@Z_TOC LVL 2 = 6.	 Operational Experience and Performance Evaluation	22

@Z_TOC LVL 3 = 6.1.	 Initial Receiver Setup and Adjustment	22

@Z_TOC LVL 3 = 6.2.	 Receiver Operation	23

@Z_TOC LVL 2 = 7.	 Future Development	26

@Z_TOC LVL 2 = 8.	 Acknowledgements	26

@Z_TOC LVL 2 = 9.	 References	27

@Z_TOC LVL 1 = Appendix A. Host Computer Interface	28

@Z_TOC LVL 2 = A.1.	 Programming Registers	28

@Z_TOC LVL 3 = A.1.1.	 0300 AMD9513A Timing Generator (control) (TGC) and 0301 AMD9513A Timing Generator (data) (TGD)	28

@Z_TOC LVL 3 = A.1.2.	 0302 A/D Converter Buffer and Multiplexor Address (ADCBUF)	28

@Z_TOC LVL 3 = A.1.3.	 0303 A/D Converter Status and A/D Converter Start (ADCGO)	29

@Z_TOC LVL 3 = A.1.4.	 0304 D/A Converter A (DACA) and 0305 D/A Converter B (DACB)	29

@Z_TOC LVL 3 = A.1.5.	 0306 Pulse-Code Buffer (CODE)	29

@Z_TOC LVL 3 = A.1.6.	 0307 Parameter Buffer (PAR)	29

@Z_TOC LVL 2 = A.2.	 Timing Generator Operation	30

@Z_TOC LVL 3 = A.2.1.	 Counter 1 (P0, P1)	30

@Z_TOC LVL 3 = A.2.2.	 Counter 2 (GRI)	31

@Z_TOC LVL 3 = A.2.3.	 Counter 3 (PCX)	31

@Z_TOC LVL 3 = A.2.4.	 Counter 4 (STB)	31

@Z_TOC LVL 3 = A.2.5.	 Counter 5 (OUT)	31

@Z_TOC LVL 2 = A.3.	 Interface Connectors	31

@Z_TOC LVL 3 = A.3.1.	 J201 (ANT) Antenna/Preamplifier Input	31

@Z_TOC LVL 3 = A.3.2.	 J203 (5MHZ) Standard Frequency Output	32

@Z_TOC LVL 3 = A.3.3.	 J204 (S1) Q-Integrator Output	32

@Z_TOC LVL 3 = A.3.4.	 J205 (PCI) Envelope Gate Output	32

@Z_TOC LVL 3 = A.3.5.	 J206 (SIG) 100-kHz Signal Output	32

@Z_TOC LVL 3 = A.3.6.	 J301 (OUT) Counter/Timer Output	32

@Z_TOC LVL 1 = Appendix B. Program Listing	33

@Z_TOC LVL 1 = Appendix C. Schematic Drawings	56

@Z_TOC TITLE = List of Figures

@Z_TOC LVL 6 = Figure 1. LORAN-C Timing Receiver	2

@Z_TOC LVL 6 = Figure 2. LORAN-C Group Repetition Interval	3

@Z_TOC LVL 6 = Figure 3. LORAN-C Pulse Coding	3

@Z_TOC LVL 6 = Figure 4. Master Autocorrelation Function	4

@Z_TOC LVL 6 = Figure 5. LORAN-C Signal	4

@Z_TOC LVL 6 = Figure 6. Receiver Block Diagram	6

@Z_TOC LVL 6 = Figure 7. Receiver Controller Block Diagram	9

@Z_TOC LVL 6 = Figure 8. Receiver Selectivity	13

@Z_TOC LVL 6 = Figure 9. Receiver Bandpass	13

@Z_TOC LVL 6 = Figure 10. Receiver Output	13

@Z_TOC LVL 6 = Figure 11. RMS Error Amplitude	22

@Z_TOC LVL 6 = Figure 12. Prototype Timing Errors	24

@Z_TOC LVL 6 = Figure 13. Prototype Frequency Errors	24

@Z_TOC LVL 6 = Figure 14. Austron Timing Errors	24

@Z_TOC LVL 6 = Figure 15. Austron Frequency Errors	24

@Z_TOC LVL 6 = Figure A.1. Receiver Controller Registers	28